BROOKHAVEN NATIONAL LABORATORY

LASER CONTROLLED AREA

STANDARD OPERATING PROCEDURE (SOP)

This document defines the safety management program for the laser system listed below.  All American National Standard Institute (ANSI) Hazard Class 3b and 4 laser systems must be documented, reviewed, and approved through use of this form.  Each system must be reviewed annually.

System description: 20 Watts CW/pulse fiber-coupled semiconductor diode laser

Location: AGS A3 beamline for E951 experiment

### LINE MANAGEMENT RESPONSIBILITIES

The Owner/Operator for this laser is listed below.  The Owner/Operator is the Line Manager of the system and must ensure that work with this laser conforms to the guidance outlined in this form.

| Owner/Operator: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Name: | Kirk McDonald (Princeton) | Signature: | Date: |  |

# AUTHORIZATION

Work with all ANSI Class 3b and 4 laser systems must be planned and documented with this form.  Laser system operators must understand and conform to the guidelines contained in this document.  This form must be completed, reviewed, and approved before laser operations begin.  The following signatures are required.

| BNL LSO printed name | Signature | Date |
| --- | --- | --- |
| ES&H Coordinator printed name | Signature | Date |

APPLICABLE LASER OPERATIONS

x General Operation x Alignment  Service/Repair  Specific Operation x Fiber Optics

ANALYZE THE LASER SYSTEM HAZARDS

Hazard analysis requires information about the laser system characteristics and the configuration of the beam distribution system.

Laser System Characteristics

| Laser Type (Argon, CO2, etc) | Wavelengths | ANSI Class | Maximum Power of Energy/Pulse | Pulse Length | Repetition Rate |
| --- | --- | --- | --- | --- | --- |
| Semiconductor diode | 808 nm | IV | 20 Watts CW/peak | Variable, 15-150 ns | CW – 1 MHz |

Cryogen Use

Describe type, quantity, and use.

none

Chemicals & Compressed Gasses

Describe type, quantity, and use.

none

Electrical Hazards

Description (Describe the power supply to the system).

CW operation – 4 volt dc with 20 Amp maximum current.

Pulsed operation – external trigger at 1 MHz repetition rate with a bunch length (gate width) of 16-microsecond. Trigger runs at one bunch per 5 minutes. Pulse width is set at 150 ns, therefore the duty cycle is 15 %.

Other Special Equipment

Description (Equipment used with the laser(s)).

A Labview interfaced PC controller for the current pulser.

Laser System Configuration: Describe the system controls (keys, switch panels, computer controls), beam path and optics (provide a functional/block diagram for complicated beam paths).

This diode laser itself does not has an ON/OFF power switch. It is driven by a current pulse source, which is remotely controlled using a Labview interface on a local PC. Because of the potential radiation hazard during beam run within the vicinity of the equipment, the operators access and control the local PC through PC-Anywhere or a Virtual-Network-Interface from a remote PC. Therefore, a remote PC actually controls the laser. The laser itself is in a sealed package, together with the heat sink it is enclosed in an aluminum housing show in Figure 1.

Figure 1

Layout of the optical beam transport in the A3 beamline for experiment E951 is shown in Figure 2 and Figure 3. The thin red line is the fiber optics line and the thick solid red line is the optical beam propagating in free space. The laser beam is fiber-coupled into the radiological controlled area where it is diverged at a full opening angle of 23 degree at the fiber output end. A lens then collimates the diverged light to an area of 6 inch in diameter. A mirror steered the optical beam to a diagnostic port near the mercury target area.  The fiber coupling terminates at the lens where the light then travels unenclosed with a spatial beam diameter of 6 inch in diameter and exits through the radiation shielded wall where it  terminates on a telescope/CCD camera arrangement. The spatial size of the laser beam is converged to 2-3 inch in diameter at the entrance aperture of the telescope. A second telescope/CCD views the window of the target continuously in real time on an off-axis optical beam path – green/yellow line, no laser light propagates in this beam path.

Figure 2

Figure 3

### DEVELOP CONTROLS

### IDENTIFY ES&H STANDARDS

Recognition, evaluation, and control of laser hazards are governed by the following documents.

American National Standards Institute (ANSI) Standard for Safe Use of Lasers;

(ANSI Z136.1-2000)

Laser Safety Subject Area

Brookhaven National Laboratory Environment Safety and Health Standard: 1.5.3 INTERLOCK SAFETY FOR PROTECTION OF PERSONNEL

ENGINEERING CONTROLS

x Beam Enclosures			x Protective Housing Interlocks		 Other

x Beam Stop or Attenuator		 Key Controls

Activation Warning System		 Other Interlocks

Ventilation				 Emission Delay

Describe each of the controls in the space provided below this text.  Interlocks and alarm systems must have a design review and must be operationally tested every six months.  Controls incorporated by the laser manufacturer may be referenced in the manuals for these devices.  Attach a copy of the design review documentation and a written testing protocol.  Attach or keep elsewhere any completed interlock testing checklists to document the testing history.

Engineering Controls Description:

This laser is of Class IV when it operates in CW mode (maximum of 15 Watts). When it is operated in the pulsed mode (peak power of 15 Watts in 150 ns pulse width at 1 Hz repetition rate), its classification is lowered to Class IIIb. The high CW power is used only during final beam alignment. To lower the laser radiation hazard during beam alignment, a low power (1 mWatt) CW red laser diode is used to align all optics prior to the final infrared beam alignment. A metal beam enclosure is placed at the location where the fiber cable couples to the free space optics where laser beam is diverged to the eye level position. A laser interlocked beam shutter is placed at the exit end of the fiber cable. The shutter is interlocked to the  second entrance gate of the A3 beamline. The location of the interlock switch and the activation button is shown in Figure 4.

Figure 4

The operation of the laser interlock is described here. During laser alignment, that is, when laser operator/s are inside the cave.

- Close the 2nd entrance gate of the A3 beamline.

- Wear laser safety goggle.

- Turn laser on (laser beam should be blocked by the beam shutter in its default mode).

- Interlock the laser controlled area in the cave by pushing SW1 once. An indicator light on SW1 will light up.

- Open the laser shutter by pushing SW3 once. An indicator light on SW3 will light up. Laser light will be emitting. Also switch SW4 will be de-activated and the condition inside the cave will be displayed on SW4 with various indicator lights.

- Upon exiting the area, opening the 2nd entrance gate or pushing SW1 once will close the beam shutter and de-activate all switches.

In preparing the laser system for the experiment run.

- Clear all personnel inside the A3 cave.

- Close the 2nd entrance gate and interlock the laser controlled area by pushing SW2 once, appropriate indicator light will light up on SW2 and SW4.

- To open the laser shutter, push SW4 once. Laser light will be emitting and a flashing laser warning sign at the vicinity of SW4 will be turned on, and appropriate indicator light will light up on SW2 and SW4.

- Upon completion of the experiment, turn laser off and push SW4 once.

ADMINISTRATIVE CONTROLS

x Laser Controlled Area	  x Signs	x Labels	 Operating Limits

The format and wording of laser signs and labels are mandated by BNL and ANSI standards.  Only the standard signs are acceptable.  Standard signs are available from the BNL Laser Safety Officer.

All lasers must have a standard label indicating the system’s wavelength, power, and ANSI hazard class.  Required labels must remain legible and attached.  The manufacturer should label commercial systems.

Standard Operating Procedures (SOP) are required for laser system operation, alignment, and maintenance.  The SOPs need only contain the steps necessary to perform these tasks and identify when and where posting and personal protective equipment is required.  SOPs must be approved by the BNL Laser Safety Officer and should be kept with this program documentation.

Administrative Controls Description:

Laser warning labels are placed at the second entrance gates of the A3 beamline. A flashing laser warning sign is placed at the vicinity of the optical setup inside the A3 cave. Similar warning label is also placed at the vicinity of the optical equipment outside the A3 beamline. The telescopes and the CCD cameras are enclosed in a metal frame. All E951 personnel are briefed about the laser hazard and the safe operation of the laser. At least one qualified laser operator should be at the control console when any one member of the E951 personnel operates the console. Laser beam alignment is done by qualified, trained, laser operators. In E951 experiment only two qualified laser operators, Thomas Tsang and Harold Kirk, are allowed to align and manipulate the optics.

There are 4 sections of optical alignment procedures.

The first procedure is proceeded from the beam pipe P1, where the telescope is located, to the fiber-optic assemble A1.

- Shine a low power (1 mW ) red laser diode at the center of the beam pipe, P1, to the final turning mirror, M3.

- Align red laser beam from M3 to M2.

- Align red laser beam from M2 to M1.

- Align red laser beam from M1 to the center of L1 and onto the center of the fiber-optic cable launch assembly, A1.

- Plug the exit end of the beam pipe, P1, with metal disk.

The second procedure is proceeded from the fiber-optic assemble A1 back to the beam pipe P1.

- Connect a fiber coupled 1 mW red laser diode to the entrance end of the jumper fiber cable F1.

- Adjust the focusing of the lens L1 until an optimum focusing condition is obtained on P1.

- Make any necessary mirror adjustment to center the beam on P1.

The third procedure is done using the IR laser operated in CW mode.

- Interlock and secure the laser-controlled area. Wear laser protective goggles.

- Disconnect the fiber-coupled red laser diode from fiber cable F1.

- Connect fiber cable F2 to F1.

- Operate the IR laser in CW mode at 7 Amp drive current (1 Watt optical power output).

- Make all necessary minor optical adjustment to send IR beam from A1 to the center of the beam pipe P1.

- Confirm the beam alignment by viewing the IR beam using an IR fluorescence card or an IR handheld viewer.

Laser alignment is done.

The last procedure is to align a second telescope for off-axis viewing the target window.

- Shine a low power (1 mW ) red laser diode at the center of the beam pipe, P2, to the turning mirror, M5.

- Align red laser beam from M5 to M4 and onto the center of the target.

All optical alignment is done.

Because the laser diode is packed in a completely sealed package, no laser maintenance or service is required. However, the optical alignment may need to be checked periodically by following the second alignment procedure with a low power red diode laser.

The optical setup of the E951 experiment is shown in Figure 5.

Figure 5

A trigger signal is derived at the arrival of the proton beam. It triggers the fast camera and a Gage CompuGen 1100 PC board pulse generator at a rate of once every 5 minutes. The Gage board generates a 16-pulse train per proton bunch. Each pulse has 0.5 microsecond duration in one microsecond period. The electrical pulse train is then used to trigger the Avtech current pulser. The operating procedures are (sample front panel is shown below):

- From the Start menu, open E951-optics menu, select Start-Electronic-trigger.

- Click Waveform Editor.

- Open signal files in D:\E951-optics directory: GaGe-16-pulse-1-us.sig for 16-pulse 1 microsecond period, GaGe-16-pulse-10-us.sig for 16-pulse 10 microsecond period, etc..

- Expand and view the signal waveform.

- Select Setting – Trigger Setting – External.

- Select Setting – Loop Setting – Multishot – 4.

- Click GEN button on the Waveform Editor panel to generate and output the waveform.

Figure 6

The laser diode is driven by an Avtech current pulser model # AVOZ-A1A-B-N. This pulser is controlled by a LabView interface via an IEEE bus and is triggered by the Gage board.  The operating procedures of the laser driver are (sample front panel is shown below):

- From the Start menu, open E951-optics menu, select Start-laser-trigger. An Avtech-pulse.vi will initiate.

- Set (or confirm) IEEE address to 8.

- Set output voltage level to –20 volt.

- Set (or confirm) trigger to External.

- Set (or confirm) pulse width to 150 ns.

- Initiate the pulser by running the vi once.

- When proton beam is available, activate the laser interlock. Click the Output Status button to LASER ON and run the vi.

- Laser is now ON and is firing 16 pulses, 150 ns duration per pulse, at the repetition rate of the proton beam – once every 5 minute.

Figure 7

A fast CCD camera, SMD-64K1M that is also triggered by a trigger pulse arriving once every 5 minutes captures the image of the target. The camera has a dynamic range of 11-12 bit (4096) of resolution. Its operating procedures are (sample front panel is shown below):

- From the Start menu, open E951-optics menu, select Start-fast-camera.

- Click Acquire (the icon with a camera) once. System will then initiated.

- Set Ext. Trig. to SYNC, Integ. Time to 0.001, gain to 1.0, and pause time to 0.

- Double click on the Acquire button. Acquire Parameter menu will pop-up.

- Check the Non-stop Acquire box. If single image viewing is preferred check the Raw Multi-Exposure box (the single image will be displayed in a nice rectangular window rather than in a tilted window).

- Go to the main display menu. Select Tool-Options-Nonstop acquire display-image statistics submenu. Uncheck the Show Histogram and Show General Statistics boxes.

- Click OK. 16-image or a single image will be displayed in real time.

- To save image: click File – Save As – check 12-bit option – type in the file name.

Figure 8

The system is now ready for the experiment. Display all front panels on the monitors and activate the VNS (virtual-Network-Computing) interface. A sample of the control/monitor panels is shown here.

Figure 9

Detail on the characteristics of the laser is given in the following. Although this laser unit has 2 identical laser diodes each capable of emitting 7 Watt of power, only one laser diode is in used (bottom laser diode). The dc current-voltage-light and the ac voltage-light characteristics are shown in Figure 10(a). At the maximum pulse voltage of –20 volts, 12 Watts of peak power is emitted. However, the light intensity temporal profile varies with drive voltages, see Figure 10(b), and is of triangular shape.

Figure 10

When a 16-pulse train triggers the laser, the intensity of the 16 laser pulses varies, see Figure 11. The light intensity drops by a factor of 2 from the first to the third pulse. This light intensity drooping is also noticeable on the captured image, see Figure 9. However, the temporal shape of all 16 pulse is maintained - triangular. The phenomenon of light intensity drooping and triangular shaped temporal profile is a result of the limited electrical bandwidth of the laser and the limited pulse duration of the voltage driver (maximum of 150 ns).

Figure 11

During the experiment if one needs to reduce the number of photons reaching the camera by reducing the pulse width or the pulse voltage of the laser driver, one should aware that either techniques produce nonlinear laser output intensity, as shown in Figure 12.

Figure 12

## CONFIGURATION CONTROL

Prepare and attach a checklist to be used for configuration control of any protective housings, beam stops, beam enclosures, and any critical optics (mirrors or lenses that could misdirect the beam and result in personnel hazard).  Include entries to ensure placement of required signs and labels and status of interlock verification. Completed checklists must be posted at the laser location.  The checklist does not have to be redone unless there has been a system modification, extended shutdown, or change of operations.

PERSONAL PROTECTIVE EQUIPMENT

x Eye Wear		 Skin Protection

Eye Wear: All laser protective eyewear must be clearly labeled with the optical density and wavelength for which protection is afforded.  Eyewear should be stored in a designated sanitary location.  Color -coding or other distinctive identification of laser protective eyewear is recommended in multi laser environments.  Eyewear must be routinely checked for cleanliness and lens surface damage. Skin Protection: For UV lasers or lasers that may generate incidental UV in excess of maximum permissible exposure (MPE), describe the nature of the hazard and the steps that will be taken to protect against the hazard.

| EYE WEAR SPECIFICATIONS |  |  |  |
| --- | --- | --- | --- |
| Laser System Eyewear #### Identification | Wavelengths | Intra-beam Optical Density | Diffuse Optical Density |
| Edmund K53-291 | 808 nm | 3-4 |  |
| Edmund K38-218 | 635 nm | >5 |  |

Define eyewear optical density requirements by calculation or manufacturer reference and list other factors considered for eyewear selection.  The BNL Laser Safety Officer will assist with any required calculations.

1. For invisible beams, eye protection against the full beam must be worn at all times unless the beam is fully enclosed.

2. For visible beams, eye protection against the full beam must be worn at all times during gross beam alignment.

3. Where hazardous diffuse reflections are possible, eye protection with an adequate Optical Density for diffuse reflections must be worn within the nominal hazard zone at all times.

4. If you need to operate the laser without wearing eye protection against all wavelengths present, explain the precautions that will be taken to prevent eye injury.

### TRAINING

LASER SAFETY TRAINING

Laser Operators must complete sufficient training to assure that they can identify and control the risks presented by the laser systems they use.  Owners/Operators and Qualified Laser Operators must complete the BNL World Wide Web based training course (BNL course #HP-IND-011).

Qualified Laser Operators must also complete system-specific orientation with the system owner/operator.  System-specific training must be documented with a checklist that includes

- Trainee name and signature

- Owner/Operator signature

- Date

- Brief list of topics covered

- Review of this program documentation

- Review of SOPs

All laser safety training must be repeated every two years.

MEDICAL SURVEILLANCE

Operators of ANSI Class 3b and 4 laser systems must complete a baseline medical eye examination prior to laser system operation.  Any qualified ophthalmologist may complete this exam.  BNL has arranged for this service from the following local physicians:

Dr. Charles Rothberg		The Ophthalmic Center			East End Eye Associates

331 East Main St.		Dr. Basilice				Dr. Sherin

Patchogue, NY 11772		3400 Nesconset Highway		669 Whiskey Road

East Setauket, NY 11733		Ridge, NY 11961

631 758-5300			631 751-2020				631 369-0777

$65 per exam			$60 per exam				$125 per exam

Personnel using physicians other than those listed must have their examination records forwarded to the BNL Occupational Medicine Clinic.

FEEDBACK AND IMPROVEMENT

Comments and suggestions for improvement should be directed to BNL-Laser Safety Officer, Chris Weilandics (X2593; weil@bnl.gov).

LASER USER QUALIFICATION

Personnel qualified to work with this laser system are listed below.  These Qualified Laser Operators must understand the information and conform to the requirements contained in this document.  For training and medical surveillance, enter the date of completion.

Qualified Laser Operators:

| Basic Laser Training | Job-Specific Training | Medical Surveillance | Printed Name | Signature | Owner/Oper. Initial/date |
| --- | --- | --- | --- | --- | --- |
| 10/25/00 | X | 2001 | Thomas Tsang |  |  |
| 4/26/99 | X | X | Harold Kirk |  |  |